## Slide 1
I08-SXM:
Experiences in iterferometric control of piezo stages
at the Diamond Light Source
Brian Nutter Trevor Bates Lee Hudson

## Slide 2
Diamond Light Source at a glance
Largest single scientific investment in UK for 45 years with £M500 on completion
33 BL with 36 (37) end stations in 2018 after completion of phase III
Funded by the UK government through STFC (86%) and the Wellcome Trust (14%)

## Slide 3
Diamond Light Source at a glance
1026 proposals awarded beamtime
9091 shifts awarded
5472 user visits
3252 unique visitors
1011 unique PhD student visitors
967 journal papers
In 2016:

## Slide 4
Imaging and spectromicroscopy facilities @ Diamond
I08 Scanning X-ray Microscopy
I14 Nanoprobe/
eBic/ ePsic
I13 X-ray Imaging and Coherence
I18 Microfocus Spectroscopy
I12 Joint Engineering, Environmental and Processing
B16 Test beamline
B24 Cryo-TXM
B22 Infrared Microspectroscopy
I05 nano-Arpes

## Slide 5
Presentation Outline
I08 soft x-ray Microscope – what it is what it does and why change
Motion Control Journey – a quick walk through the early stages
Motion Control Journey – a quick look at the current results
Motion Control Journey – what challenges were overcome
I08 Upgrades – motion control in perspective
Trajectory scanning at Diamond – what we do
Trajectory scanning at Diamond – some problems to overcome
Credits
Conclusions

## Slide 6
I08-SXM: The Diamond scanning X-ray microscopy BL
Broad photon energy range (250 – 4200 eV)
Highest resolutions down to 20 nm for some imaging modes
up to 1010 ph/s in a 20nm spot
Various detection modes for imaging and spectro-microscopy (XANES, LEXRF, electron emission)
Versatile specimen environments including cryogenic (LN2) specimen environments
Principal branch
Microprobe forming ZP
Specimen raster-scanned across the microprobe
Transmission detector system (EMCCD)
Versatile specimen environments (with focus on cryo)
Secondary photon or electron detectors in back-scattering configuration
Notes: zone plate focuses the beam through an aperture to a spot size down to 20Nm.
The sample in its environment is scanned across the beam to build up a picture
The detector is a photon counting detectror. Two other detectors can take data simultaneously

## Slide 7
I08-SXM: End station
Vacuum
window
ZP stage with scanner
XRF detector
OSA
Laser interferometer
Transmission detector system
Specimen coarse and fine stage
Cryo stage
Cold trap

## Slide 8
I08-SXM: End station

## Slide 9
The I08 SXM: upgrade of the controls and DAQ system
I08 is a commercial end station with commercial DAQ hardware and software with a ‘lack of control on the controls system’:
Scanning Project
I08 optimised project for improving scanning hardware and interferometric control
Mapping Project
Generic Diamond project for common scanning interfaces and data analysis software
Faster data acquisition
Different scan profiles
Better synchronisation of multiple detector geometry
Online processing
Simplify support and maintenance
Everyone is happy
and we are all in
good mood!

## Slide 10
A Motion Control Journey: Sample X
Diamond Light Source continues to develop and upgrade ‘Modules’ to keep pace with the latest available technology. Beamline I08 recently upgraded the end-station using the Attocube IDS3010 interferometer.
It started as a development rig simulating the three orthogonal sample positioning systems, X Y and Z. Consider the X Axis (horizontally across the beam).

## Slide 11
A Motion Control Journey : First Noise
First impressions reading the AQuadB differential feedback TTL signals showed a lot of noise - To be expected at this resolution. (Disregard the scaling at this early stage)

## Slide 12
The Motion Control Journey : A few problems - RTFM
The A-quad-B process is a digital interface that allows high resolution over large sizes of data. Resolution and Clock define the quadrature process. The signal is encoded over two different 2-level-channels, A and B. The levels’ amplitude is either LVTTL or LVDS (one at a time):
 Single-ended with LVTTL levels
 Differential with LVDS levels

## Slide 13
The Motion Control Journey: BISS-C to the rescue?
Maximum A Quad B = 6 MHz @1nm resolution = 6mm per second
We want finer resolution - to improve spacial resolution
We want higher speed – to keep up with coarse measurements
IDS3010 can get down to 1pm resolution and up to 20MHz
BissC – a serial absolute encoder format .
IDS3010 – Provides a 32 bit window onto its own internal 48 bit position register. The selected LSB determines the resolution. Minimum 1pm.
Further there are 2 Error bits + 6 CRC bits
GeoBrick uses a 2MHz clock to sample the data , updates each axis at 5KHz
Possibility to combine AquadB of stepper X with BissC of piezo X in 1 plug.

## Slide 14
The Motion Control Journey: The Proposed Beamline Set Up.
PandA is the new name for ZEBRA-2 . Takes BissC .

## Slide 15
The Motion Control Journey: BISS-C to the rescue?
As supplied
Sanitised – becomesC1
AquadB plus BissC is C4
PandA box inserted between C4/2 and C3/1 etc

## Slide 16
The Motion Control Journey: And the Piezo?
Diamond standard motion controller for Phase 2 beamlines. 8 on board amplifiers. 8 incremental encoders. Optional Fibre Optic MACRO interface. Optional BissC Interface
Acc24M2A 2 channel MACRO interace with 2 channels differential +/- 10 V control outputs 16 bit resolution. 2 channels incremental encoder
PI E505 piezo amplifiers. 3 axis integrated flexure stage with 200um x 200um x 20um travel. Cap sensor feedback.

## Slide 17
The Motion Control Journey: Noise Measurements
Put the test rig on a table and see what you get. Answer the question is it any good?
135Hz and 178Hz peaks

## Slide 18
The Motion Control Journey: Noise Measurements
Then put it on an optical bench and see the noise reduce by factor of 10. Answer yes its worth a try

## Slide 19
The Motion Control Journey: In Situ Performance Tests
Axis 1-X Coarse -Stepper motor moves 500 counts of 10nm resolution.
Axis 9-X Fine -Interferometer measures 50000 counts of 100pm resolution. But we see the effect of the cogging in the stepper.
Movement of Coarse X measured by open loop Fine x
Notes: The distutrbances in the blue line. Are they cased by Abbe errors On Axis?

## Slide 20
The Motion Control Journey: In Situ Performance Tests
Axis 1-X Coarse -Stepper motor moves 500 counts of 10nm resolution.
Axis 9 –X Fine attempts to maintain its own zero position, reversing the motion of Coarse X.
Here Fine X counters the commanded motion of Coarse X

## Slide 21
The Motion Control Journey: In Situ Performance Tests
The effect of an exponential filter on the readback of the fine encoder
The piezo is driven closed loop on axis 5.
Move is 25000 counts in 10ms. We see some overshoot and also some noise when ‘in-position’.
Axis 8 is reporting the same interferometer as 5, but its result is passed through an exponential filter – eliminating feedback noise at the expense of a lag.
Notes: What does the scientist want. High Speed Dynamic response, or smooth control with less dynamic response

## Slide 22
The Motion Control Journey: In Situ Performance Tests
10 steps of 500 increments in Fine X. Observed FE is a function of the filter.

## Slide 23
The Motion Control Journey: In Situ Performance Tests
When we move fine X what is happening to fine Y and fine Z

## Slide 24
The Motion Control Journey: In Situ Performance Tests
It seems as though there is some orthogonality issues, especially in Z . This could be mounting error of either the piezo stage unit, or the interferometer mirrors. We can either use compensation tables or include y and Z offsets in the X axis definition

## Slide 25
What were the Motion Control Challenges
Why should we replace a functioning system ?
We have ongoing support and reliability issues for XPS
This is really 3 separate systems on 1 industrial pc – obsolete upon purchase
To integrate it into the mapping project linking motion and data collection
Can our motion controller do it ? What were we trying for the first time?
Closing the position loop around an analogue output of position over MACRO
using a PID loop designed for positioning via a torque demand
Using BissC prototype from the interferometer supplier
Using two motors on top of each other to position in 1 DOF
coarse-coarse moves with Piezo reducing settling times
fine only moves for rapid scanning over piezo range
coarse – fine – coarse for high resolution scans over extended range
Develop ‘custom servo’ programs to insert ‘position feedforward’ into loop

## Slide 26
What were the Motion Control Challenges
How is it possible to do it without interrupting the beamline?
As much proof of concept testing as possible offline
In a suitable environment
Initial in situ testing using the original cabling with adaptor boxes
this allows us to ‘put it back for users’ and beamline commissioning
at the expense of introducing noise through leaky connections
Testing only for short intermittent periods on machine days and shutdowns
What else is happening at the same time?
New motion programme development to link to mapping project
pushing processor and memory performance to its limits
2 other beamlines awaiting developments to use same technology from new
2 other beamlines watching to see how ‘in-line’ piezos can improve performance

## Slide 27
The I08 upgrade of the controls, DAQ and analysis system
Scanning and mapping DAQ and analysis workflow:
GDA (Generic data acquisition) or user interface
Key contributors
of a team of >15 engineers/ developers:
Malcolm middle layer with python controlled hardware box
EPICS interfaces
Savu scripts for online processing
DAWN for data analysis/ 3rd party analysis software
Motion controls
Detectors
Saving data in HDR SWMR files
Mark
Basham
Tom
Cobb
Jacob
Filik
Aaron Parsons

## Slide 28
Trajectory Scanning on the Motion Controller
Motion program on PMAC executes generic PVT moves
Position and time coordinates, as well as user subroutines, are read from user memory
Depending upon memory allocation Buffers A and B may contain up to 3000 sets of positions

## Slide 29
Trajectory Scanning on the Motion Controller
Each move definition has a 48 bit word for each axis
(A, B, C, U, V, W,X, Y, Z ) using X and Y memory
Delta time, a user subroutine (if any) and a velocity mode are provided in a separate sub-buffer
Most common subroutine is to trigger i/o at segment boundaries to give triggers to the PandA card

## Slide 30
Motion
Program
Motor
Demand
Support for Turbo PMAC and Power PMAC
Move raw and co-ordinate system motors
PVT scanning of motors
Double buffered points sent to motion program
pmac
driver
Positions + Build
Run
Positions + Append
Trajectory Scanning on the Motion Controller

## Slide 31
Trajectory Scanning on the Motion Controller
The Mode bits select which data points to use to calculate velocity in te PVT calculation
The 4 types are
Previous to current
Previous to next, which is the normal mode of operation
Current to next
And zero velocity, useful for commanding a dwell to allow other actions to take place

## Slide 32
Trajectory Scanning on the Motion Controller

## Slide 33
Trajectory Scanning on the Motion Controller
Dropped Frames or Slow Motion
  Approximately 1 in 1000 segments would drop a frame or run slow ???
Diamond standard servo update rate is approximately 5 KHz
  PVT actions update at the servo ‘tick’ for best repeatability
  Scanning time is defined in units of ¼ millisecond
  Velocity calculations are made with a predetermined millisecond calculator
  Turbo PMAC is not a true floating point controller
First fix: select velocities and segment times that never skip a frame
  But this resulted in approx 1/1000 segments going slower than expected
Current fix: a3.2KHz servo update rate has been applied in other installations
  This gives an integer number of ¼ microsecond ticks and an integer millisecond calculator. There has been no major drop off in perfromance of the stepper axes but it may require signifiant retuning of the other axes.

## Slide 34
The I08 upgrade of the controls, DAQ and and analysis system
Benefits of the Mapping and Scanning projects:
Significant reduction of time overheads in the scans, and therefore radiation dose
Fast on-the fly large area scans with combined stepper/ piezo motors and interferometer feedback:
Arbitrary scan profiles (spiral scans, trapezoids, rectangular scan with offsets, etc)
Online PyMCA fitting of XRF spectra and display of fitted data during a raster scan (available to users at I08)
Online reconstruction of ptychographic data sets and display of reconstructed images during a raster scan (work in progress at I08, available to users soon)

## Slide 35
Jon Kelly,
Mechanical
Engineer
Leo Cahill,
Electrical Engineer
Nic Bricknell
Software System Engineer (Controls)
Eva G. Navarro,
Detector Scientist
Chris Obin
Project Planning Officer
Hugo Shiers,
Vacuum Engineer
Hongchang Wang,
Optics Scientist
Matthew Hand,
Optics Scientist
Chris Bloomer
Diagnostics Scientist
Anthony Hull
Software Engineer (Data acquisition)
Jacob Filik
Software Engineer (Data analysis)
Trevor Bates
Motion Controls Engineer
Lee Hudson
Motion Controls Engineer
Alex Bugnar
Nanopositioning and metrology
CREDITS – I08 SXM working group

## Slide 36
Conclusions
Has it been a huge success?
  It is widely acknowledged that by converting from XPS controller to the Delta Tau platform we have significantly improved the performance of the SXM
  It has taken a lot of effort to make the low level hardware work, the challenges have been electrical, mechanical and developing new ways of making the old controller tick.
  What started out as a beamline specific scanning upgrade project has become an integral test case for the wider mapping project. Whilst the I08 SXM has some unique challenges it has ebnabled similar techniques to be used in 3 other beamlines all intvested with the mapping project?
  A new side branch for I08 will use similar techniques and by using ptychography will enable spacial resolutions far below 10nm.

## Slide 37
Questions.
How much is a standard worth?
  Do we try to get tomorrows performance out of yesterdays controller?
ANY QUESTIONS?
And may I draw your attention to …
TUPHA159: Malcolm: A Middlelayer Framework for Generic Continuous Scanning
THPHA068 PandA Blocks Open FPGA Framework and Web Stack
THPHA200 BART: Development of a Sample Exchange System for MX Beamlines
THPHA101 Review of Personnel Safety Systems at DLS
TUPHA212 Odin - a Control and Data Acquisition Framework for Excalibur 1M and 3M
THPHA115: A New Transverse and Longitudinal Bunch by Bunch Feedback Processor

## Slide 38
Thanks for your attention!

## Slide 39
Spatial resolution
Photon energy
Gray zone
Analysis from the nanoscopic to macroscopic scale

## Slide 40
J08: End station layout
In-house development with various components to be subcontracted as modules
Sample changer mechanism
Interferometric control of sample positions and optics
Multi-detector geometry (transmission CCD/ XRF detector)
Cryo rotation stage
Controls and data acquisition software is a 1:1 copy from the main branch (Mapping and Scanning projects)
Conceptual design:

## Slide 41
J08: End station layout
Andor Balor camera
Cryo-rotation sample module
ZP/ OSA module
XRF detector
Sample
changer
Leica transfer
system

## Slide 42
I08-SXM: End station
Vacuum
window
ZP stage with scanner
XRF detector
OSA
Laser interferometer
Transmission detector system
Specimen coarse and fine stage
Cryo stage
Cold trap

## Slide 43
J08: End station layout
Cryo-cooled
Sample ‘hotel’
Cryo-rotation sample module
Sample
gripper
Leica transfer
system
Sample handling system:
(inspired by I14 designs)

## Slide 44
J08: End station layout
Sample dome:
Cryo-cooled interface based on I23 goniometer
Cold nitrogen gas cooling as at I08
Ready to build a prototype
Sample holder compatible to I14 and EMs
Piezo scanner
Stepper x stage
Stepper elevation
stage
Rotation stage
Cryo-cooled interface
Stepper actuators recovered from Scanning Project

## Slide 45
J08: End station layout
Sample dome:
Cryo-cooled interface based on I23 goniometer
Cold nitrogen gas cooling as at I08
Ready to build a prototype
Sample holder compatible to I14 and EMs
Sample
Interferometer mirror
Interface plate with active heaters and thermo sensors
Rotation stage
Cold stage

## Slide 46
Imaging and spectromicroscopy facilities @ Diamond
I08 Scanning X-ray Microscopy
I14 Nanoprobe
I13 X-ray Imaging and Coherence
I18 Microfocus Spectroscopy
I12 Joint Engineering, Environmental and Processing
B16 Test beamline
B24 Cryo-TXM
B22 Infrared Microspectroscopy

## Slide 47
J08: Spectro- and tomo-ptychography
250 - 2000 eV photon energy range
Resolutions far below 10 nm, resolution-limited only by the wavelength
Absorption and phase spectroscopies on nanoscopic scales
Cryo-Tomography
Side branch for soft X-ray diffraction imaging (Ptychography)
Operational in Summer 2019

## Slide 48
I08-SXM: Optical layout

## Slide 49
I08-SXM BL layout
Control Rooms 1&2/ Specimen Prep Area
(±0.5 oC; <40 dB;
Office environment)
Experiments Cabin
(±0.1 oC; <40 dB;
Lab environment)
Optics Cabin II
(±0.5 oC; <45 dB;
Optics Cabin I
(±0.5 oC; <45 dB)
Optics Hutch
(±0.5 oC; <50 dB)
Airlock
Vestibule
Specimen prep
lab

## Slide 50
I08-SXM: Some examples
XRF mapping of bacteria in Fe containing sediments:
Courtesy of George Flynn (New York State University/ NASA) and Sue Wirick (Chicago University/ NSLSII)
C
Fe
O
2 mm
2 mm
C
Fe
N
Courtesy of Helen Downie and Vicky Coker, Manchester University
Measurement under LN2 cryogenic conditions (-160 deg C)

## Slide 51
Insert Title Here