## Slide 1: Predictive Modelling and Error Budget Experience with HD-DCM Project
MOCRAF
Hardware challenge session
Ricardo Caliari – SIL
ricardo.caliari@lnls.br
Gabriel Moreno – SIL
gabriel.moreno@lnls.br

## Slide 2: Outline
High-Dynamic DCM concept
Predictive modelling
Error budget
System characterization
Partial results

## Slide 3: Motivation
DCM performances for the next generation machines
10%
source
size
Sirius

## Slide 4: Specifications
Parameter | Description
Type: | Vertical DCM
Beam offset: | 18 mm
Angular range: | 3 to 60°
Angular resolution: | 0.2 μrad
Pitch/roll stability: | in-position: < 10 nrad (+3σ) flyscan: 200 nrad (+3σ)
Crystal sets: | Si(111): 2.3 to 38 keV Si(311): 4.4 to 72 keV
Crystal sizes (W x L): | 1st crystal: 15 x 35 mm2 2nd crystal: 15 x 190 mm2
Crystal cooling: | 1st crystal: Indirect LN2 (80 K) 2nd crystal: Copper straps (155 K)
Crystal DoF: | 1st crystal: fixed at rotation center 2nd crystal: gap, pitch, roll
Beam size: | 1.7 x 1.7 mm2
Input Power: | 150 W
Base pressure: | < 5 x 10-8 mbar

## Slide 5: High-Dynamic DCM
PCT/BR2017/050262, Patent Application.

## Slide 6: Concept
Schematics
1D dynamic model

## Slide 7: Predictive Modelling
12/2015
Voice Coil – Balance mass
Concept presentation
09/2016
Stiffness definition
Detail mechanical project
11/2016
Robustness evaluation
Parameters uncertainty evaluation
03/2017
System identification
Target achieved
Robustness test made
Lumped-mass model
and Error budget
1D dynamic model
6DoF model
System identification

## Slide 8: Error Sources to Budget
Alignment Errors (Mounting)
Thermal Drifts
IFM Mirror Surface Flatness
Cooling System Vibrations
Floor Vibrations
Instrument Induced Errors
  Amplifiers
  Actuator Disturbances
  Sensor Disturbances (i.e. SDE, non-linearity, quantization...)
And Much More!!!

## Slide 9: Error budget analysis (ShS-3DoF)

## Slide 10: Example: BW and sample freq.
Analysis made in preliminary phase, to choose the minimum suitable BW
BW = 200 Hz
BW = 250 Hz
BW = 150 Hz

## Slide 11: Closed loop control
Motion control axes:
  Goniometer (1 DoF)
  Long Stroke (1 DoF)
  Short Stroke (3 DoF)
  All linear drivers for all motors, controlled by the same ECU (same error spectra, same clock etc..)
1
2
3

## Slide 12: Short Stroke closed loop control
1
2
3

## Slide 13: Short Stroke disturbances
1
2
3

## Slide 14: Partial Results: In-Position Stability
[Kristiansen,2015]
HD-DCM Partial Results
Relative pitch: 9.2 nrad RMS
* Note: This performance is also valid for step-by-step scans, provided that the settling time is allowed.

## Slide 15: Partial Results: Flyscan w/ Stepper
HD-DCM Partial Results:
Relative pitch: < 60 nrad RMS

## Slide 16: Partial Results: Flyscan w/ Servo
HD-DCM Partial Results:
Relative pitch: < 25 nrad RMS
Problems: Servo UHV suppliers (Faulhaber, Wittenstein)
Power dissipation.

## Slide 17: See our work at ICALEPCS!
TUSH203
Speakers’ corner Tuesday – 18:15 to 18:45
THPHA214
Poster Session Thursday 16:45 to 18:20
THPHA215
Poster Session Thursday 16:45 to 18:20
Beamlines Control Platform
Rapid Control Prototyping
System Identification

## Slide 18
Thank you!
23.08.2017