## Slide 1: Sample Changer Development for Sample Environment Rise of the Robots?
MOCRAF 2017
Matt North, STFC, ISIS
matt.north@stfc.ac.uk

## Slide 2
Sample Changers
Provides sample automation.
c. 30 mins minimum to obtains useful statistics from sample for powder diffraction measurement c. 3hrs for neutron spectroscopy measurement.
In general, sample throughput, not a major cause for concern, caveat cryo experiments.
Provides ‘quick check’ for sample / technique.
Gives experimental flexibility at low cost (time)
Reduces ‘lost time’ risk across a sample set; multiplexing.
Minimise operating stresses.
Reduces support requirements.
Supports ‘Xpress’ facility access.

## Slide 3: Xpress Access
‘Xpress’ Facility access is a remote access mechanism not requiring approval from Facility Access Panel.
General Guidelines:
Short measurements c. <1 hour.
Initial sample characterisation.
Demonstrate feasibility for further proposals
Small ‘top-up’ data to complete analysis.
In the year 16/17, 25% of facility proposals were Xpress Access
Notes: Foundation of xpress on diffraction, predominantly with some SANS up to 2011.
2008 TS2 first neutrons
2009 ts2 phase 1 instruments (7) in scientific commisisoning / user programme
2011 onwards xpress service being made available to user community and broadened to include spectroscopy.
Need for spectroscopy sample changers lags – demand led?
SANS increase beyond 2008 due to sans2d and reflectometer instruments.
General increase in variety, not all supported by sample changers.
More time scheduled for xpress measurements – facilitiating blue sky in-house research

## Slide 4: Xpress Access Growth
Notes: Foundation of xpress on diffraction, predominantly with some SANS up to 2011.
2008 TS2 first neutrons
2009 ts2 phase 1 instruments (7) in scientific commisisoning / user programme
2011 onwards xpress service being made available to user community and broadened to include spectroscopy.
Need for spectroscopy sample changers lags – demand led?
SANS increase beyond 2008 due to sans2d and reflectometer instruments.
General increase in variety, not all supported by sample changers.
More time scheduled for xpress measurements – facilitiating blue sky in-house research

## Slide 5: Diffraction Sample Changers

## Slide 6: Spectroscopy Sample Changers

## Slide 7: SANS & Reflectometry Sample Changers
Notes: None of these SC’s change sample without a supporting sample positioner.

## Slide 8
SANS & Reflectometry Sample Changers
Notes: None of these SC’s change sample without a supporting sample positioner.

## Slide 9
Type | Instrument | Sample Position | Supporting Motion | Sample Container | Temp Range | Capacity | Approx Max Throughput | Number of Axis | Operating Pressures | Supporting Sample Environment
Carousel | GEM, HRPD, POLARIS | Tank | None | Cylindrical Cans; 6,8 & 11mm | 300°K | 20 samples | 10 Hours | 1x Linear 1x Rotation | 1e-2mBar to Atmos | None
Array | NIMROD, SANDALS | Tank | None | Flat Can | 300°K to 365°K | 15 samples | 15 Hours | 1x Linear 1x Rotation | 1e-2mBar to Atmos | Water Bath
Chainsaw | TOSCA | Tank | None | Flat Can | 10 - 20°K | 36 samples | 4.5 Days | 1x Rotation | 1e-6mBar to Atmos | TOSCA Cryostat
Stacked Stick | MERLIN, IRIS | Tank | None | Annular Can | 7°K to 300°K | 3 samples | 15 Hours | 1x Linear | 1e-6mBar to Atmos | Top loading CCR
Rotating MKI | SANS2D, LOQ | Open | SPS | Standard & Banjo Cuvettes | 300°K | 7 samples | 3.5 Hours | 1 rotation | Atmos | Water Bath
Rotating MkII | SANS2D, LARMOR | Open | SPS | Banjo Cuvette | 300°K to 430°K | 4 samples | 2 Hours | 4 Rotation | Atmos | Water Bath
Single Tier | LOQ | Open | SPS | Standard, Banjo & Tank Cuvettes | 300°K to 550°K | 20 samples | ~10 Hours | 0 | Atmos | Water Bath
Double Tier | SANS2D, LARMOR | Open | SPS | Standard, Banjo & Tank Cuvettes | 300°K to 550°K | 40 samples | 7 Hours | 0 | Atmos | Water Bath
7 Sample Liquid Trough | SANS2D | Open | SPS | Trough | 300°K to 380°K | 7 samples | 2.5 Hours | 0 | Atmos | None
ISIS Sample Changer Overview

## Slide 10
Type | Instrument | Sample Position | Supporting Motion | Sample Container | Temp Range | Capacity | Approx Max Throughput | Number of Axis | Operating Pressures | Supporting Sample Environment
Carousel | GEM, HRPD, POLARIS | Tank | None | Cylindrical Cans; 6,8 & 11mm | 300°K | 20 samples | 10 Hours | 1x Linear 1x Rotation | 1e-2mBar to Atmos | None
Array | NIMROD, SANDALS | Tank | None | Flat Can | 300°K to 365°K | 15 samples | 15 Hours | 1x Linear 1x Rotation | 1e-2mBar to Atmos | Water Bath
Chainsaw | TOSCA | Tank | None | Flat Can | 10 - 20°K | 36 samples | 4.5 Days | 1x Rotation | 1e-6mBar to Atmos | TOSCA Cryostat
Stacked Stick | MERLIN, IRIS | Tank | None | Annular Can | 7°K to 300°K | 3 samples | 15 Hours | 1x Linear | 1e-6mBar to Atmos | Top loading CCR
Rotating MKI | SANS2D, LOQ | Open | SPS | Standard & Banjo Cuvettes | 300°K | 7 samples | 3.5 Hours | 1 rotation | Atmos | Water Bath
Rotating MkII | SANS2D, LARMOR | Open | SPS | Banjo Cuvette | 300°K to 430°K | 4 samples | 2 Hours | 4 Rotation | Atmos | Water Bath
Single Tier | LOQ | Open | SPS | Standard, Banjo & Tank Cuvettes | 300°K to 550°K | 20 samples | ~10 Hours | 0 | Atmos | Water Bath
Double Tier | SANS2D, LARMOR | Open | SPS | Standard, Banjo & Tank Cuvettes | 300°K to 550°K | 40 samples | 7 Hours | 0 | Atmos | Water Bath
7 Sample Liquid Trough | SANS2D | Open | SPS | Trough | 300°K to 380°K | 7 samples | 2.5 Hours | 0 | Atmos | None
ISIS Sample Changer Overview


## Slide 11
CCR Adaption
Isolating VAT Valve
He Port
He Capillary & Inlet
Sensor Ports
Sample Carrier & Thermometry Contacts
Notes: Introduce gas sample change process

## Slide 12
Robot Requirements
Development Requirements:
‘Perform sample pick & place action’
‘Low cost’
‘Minimal learning curve’
‘Integrate with beamline control computer’
‘Portable package– Deliver to a beamline’

## Slide 13
Universal Robots – UR5
Motion Specification:
≥6 Axis of Motion 
Spherical Motion Envelope; 0.5m≤ r ≤1.5m 0.85m
Motion Repeatability ≤±0.1mm 
Max Payload = 5kg
Control Specification:
Local & remote control Ethernet & Pendant
Integrate with existing instrumentation A&D I/O
Motion Path Planning via 3D environment Pendant HMI
Teach Mode 
Collaborative operation (ISO 10218-1) 
Technical Specification:
‘Maintenance Free’ Operation 35,000hrs
240V Power Source
Weigh ≤20kg 18.4kg
Notes: Collaborative robot – teach mode, motion path simulation, simple programming interface, certified, with approiate risk assessment for collaborative operation.
Underlining principle – this should be delivered to sample position like any other SE.

## Slide 14
Cryogenic Robot Sample Changer
Notes: Robot test stand.
Next step to integrate robot with CCR to perform pick and place.

## Slide 15
Cryogenic Robot Sample Changer
‘Current status of the ISIS robotic cryogenic sample changer’ Journal of Neutron Research 19 (2017)
DOI 10.3233/JNR-170049
Notes: Robot test stand.
Next step to integrate robot with CCR to perform pick and place.

## Slide 16: IMAT Camera Positioning System

## Slide 18: IMAT CPS – SAT Results
Repeatability – 31.49µm (with camera exchange)
Accuracy (Pose) - 65µm
1h Holding Stability – 14.47µm
16h Holding Stability* - 45µm
Notes: *Not a requirement of the original specification
Results include robot and linear translation.
Robot system costs inc. installation and on-site commissioning c. 120k ex VAT
Does the robot engage a brake when in holding position? Are the motor drives switched of to reduce any heating effects?

## Slide 19: Future Robotic Opportunities
Science Leads
Sample Positioning – Engineering Diffraction
Texture Measurements – Powder Diffraction
Automation for Industrial QA – Reflectometry
‘The development of more automated sample environment and procedures to broaden the user base’ - Spectroscopy
Engineering Leads
Cryogenic Sample Changer - Diffraction & Spectroscopy
‘Universal’ Sample Changer - Diffraction & Spectroscopy
Notes: ‘Universal’ Sample changer for diffraction – advantages:
Increases capacity, not limited by tank volume
Provides a mechanism for elevated temperatures
Disadvantages:
System complexity
SPS+SC Combined – advantages:
Automation of varying sample environments.
Reduction in sample changer variety.
Flexibility in motion, sample environments & sample shapes & sizes.
Disadvantages:
System complexity
Space overhead.
Costly.

## Slide 20: Future Development Landscape
Funding from ISIS science group budgets for Robot Sample Changer development is not a high priority.
No step change in flux in the near term, sample throughput not a fundamental bottleneck.
Popularity of Xpress Access; growth or plateau?
Notes: No smoking gun

## Slide 21: Important steps to shift perspective
Development of in-house knowledge and skills
Demonstrating Benefit
Capitalising
Proving Capability

## Slide 22: Rise of the Robots?